- OBJECTIVE:

To describe the procedure of Calibration of Conductivity Meter.

- SCOPE:

To provide written procedure for calibration of Conductivity Meter

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Quality Control Analyst: must follow the procedure as per the SOP for calibration of the conductivity meter and perform documentation related to the operation, calibration and maintenance in the respective log.

- Quality Control Head: must train all the concerned QC Analyst before the implementation of the SOP, ensure the operation and calibration of the instrument as per SOP and ensure proper documentation

- Quality Assurance Manager approve the SOP

- ABBREVIATION AND DEFINITIONS

- SOP – Standard Operating Procedure

- KCl : Potassium Chloride

DEFINITION

- Conductivity : it is the ability of a solution to conduct electricity and the electricity is conducted in solution by ions present unit of measurement of conductivity is Siemen(S)

- PROCEDURES

# CLEANING PROCEDURE

- Switch OFF the mains power unto the Instrument.

- Use dry lint cloth to clean the Conductivity meter daily.

- Clean the electrode after each use with distilled water.

- Store the electrode in cooled MCDP water after rinsing thoroughly.

CALIBRATION PROCEDURE

- While calibration and analysis, ensure the temperature of the standard solution or sample within 25+2oC temperature

- Always use the freshly prepared KCl solution.

- Prepare standard solution of Potassium Chloride by dissolving 0.0149gm in 1000ml of distilled water.

- Rinse the electrode several times with WFI/MCDP water.

- Rinse the electrode at least twice with the Potassium Chloride standard solution to be used for the calibration.

- Dip the rinsed electrode in the standard solution and measure the resistivity of the solution at 25.0 + 20C turning the panel screw to 26.6 micro siemens. and ensure there is no air bubble in the cell and the solution is above the cell tips

- Ensure the instrument is in conductivity mode

- First press CAL and then enter and entered the cell constant of the electrode  i.e 0.01M and again press ENTER keys. The instrument will calibrate itself and shows conductivity reading followed by the cell constant value.

- Rinse the electrode thoroughly with distilled water.

- REFERENCE AND ENCLOSURE

- Operation manual supplied by the Manufacturer

- ENCLOSURE 1- WEIGHING BALANCE CALIBRATION LOG - BML-QDG-QA-036-F01

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | All Quality Control Analyst |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |

ENCLOSURE 1

Conductivity Meter Calibration Log

Doc. No.; BML/QDG/QA/037/F01

|  | Biomedical Limited, No. 1 Ohimege Road, Industrial Estate, Ilorin, Kwara State, PMB 1449 | Doc. No.: BML/QDF/QA/037 |
| --- | --- | --- |
|  | QUALITY ASSURANCE DEPARTMENT EQUIPMENT CALIBRATION LOG CONDUCTIVITY METER | Version No.: 02 |

| Date | Solution | Cell Constant | Remarks | Calibrated by | Checked by |
| --- | --- | --- | --- | --- | --- |
|  | 0.01M KCL |  |  |  |  |

Q.A Manager

Sign /Date