- OBJECTIVE:

This SOP is prepared to have a laid down standard Procedure for Calibration of glass ware

- SCOPE:

To provide written procedure for calibration of glassware

- RESPONSIBILITIES/ACCOUNTABILITIES:

Quality Control Analyst

- Preparation of the SOP

- To receive and calibrate the glassware as per procedure.

QC Head

- Review of the SOP

- Impact training to all concerned QC Analysts before implementation of SOP.

- Ensuring proper maintenance of respective documents

- ABBREVIATION AND DEFINITIONS

cGMP – Current Manufacturing practice

SOP – Standard operating Procedure

PPE – Personal Protective Equipment

wt/wt = weight by weight

wt/ml = weight per milliter

0C	- Degree Centigrade

Gm - Gram

ml	- Milliliter

- PROCEDURES

Class ‘ A’ Apparatus

- These are procured along with calibration certificate.

- On receipt calibrate the glassware before using.

- Use in work of the high Accuracy.

Class ‘ B’ Apparatus

- On receipt, calibrate it before using.

- Use in routine Work.

- All volumetric apparatus are normally calibrated at 27°C.

- Pharmacopoeia assays involving volumetric measurements require the use of accurately calibrated glassware.

- On receipt of glassware allot serial number, label accordingly and then calibrate it.

- Accept only those glassware which passes in calibration and reject those which not pass in calibration.

CALIBRATION:

- Clean the glassware & make it free from grease/dirt  before carrying out calibration

- Clean the wet outer surface with tissue paper.

- Ensure that the glassware is properly dried in the oven.

- Before calibration keep all glassware at room temp for 1 hr.

- The glassware must not be wetted above the meniscus, if this happens, wipe it with tissue paper.

- Always keep your eye at the same level, while aligning the meniscus.

- On receipt of glassware from the supplier, enter detail into the inward register with following detail

          - Date of receipt

          - Description

          - Serial No.

          - Quantity received

          - Supplier Name

          - Challan No.

          - Qty. Approved/Serial No

          - Qty rejected/Serial No.

          - Quantity received

          - Remark

- On receipt of glassware allot serial no. as per point 5.11 and then calibrate it and accept only the glassware which passes in calibration & rejects those which does not pass in calibration. Destroy the rejected glassware immediately.

- Calibrate the glassware as per attached Enclosure. The value should fall within the tolerance limits mention in the Enclosure.

For Volumetric flask

- First take the empty weight of dry volumetric flask and record the weight.

- Fill it with distilled water through funnel 5 mm below the ring mark.

- Adjust the meniscus precisely to the ring mark by adding water dropwise with a pipette.

- In case of volume exceed the marking, remove it with help of tissue paper or pipette.

- Measure the weight of the filled volumetric flask and record the weight.

For Pipette

- First take the weight of the dry weighing vessel and record the weight.

- Pipette water at approximately 5 mm above the top mark using a pipetting aid.

- Adjust the meniscus precisely by discharging the excess water.

- Allow the water to run off into the weighing vessel. So as to touch the pipette of the vessel.

- Wait for few seconds and measure the weight of vessel & record it.

For Burette

- First take the weight of the dry weighing vessel and record the weight.

- Clamp the burette vertically.

- Fill the burette to approximately 5 mm above the zero mark

- Remove any air bubble from burette & also from stop cork.

- Set the zero precisely by releasing liquid.

- Run off the water from burette as per requirement

- Ensure that burette tip shall not touch the wall of the vessel

- Wait for few seconds and measure the weight of the vessel and record it.

- Calibrate the glassware by wt/wt and convert to ml using std wt/ml at temp of water at the time of calibration referring Enclosure.

- Record all values & finding and for graduated pipettes Enclosure.

- All values and findings of graduated pipette and Burette in Enclosure.

Materials, Equipment, Apparatus and Reagents:

- Weighing balance

- Grade A glassware

- Class B apparatus

- REFERENCE AND ENCLOSURE

ENCLOSURE 1- CALIBRATION OF GLASS APPARATUS – BML-QDG-QA-

ENCLOSURE 2 - GLASSWARE CALIBRATION RECORD- BML-QDG-QA-039-F01

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | Quality Control Analyst |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |

Enclosure 1

CALIBRATION OF GLASS APPRATUS

| Name of Glassware: | Specified Capacity: | Date of Calibration |
| --- | --- | --- |
| Class: | Identification No. | Bal ID No.: |

Density of Water ( C):

| Weight of empty density bottle in gm (P) | Density of the bottle + water in Gram (Q) | Weight of water in gram ( Q-P) | Density (C) = (Q-P) Volume in ml |
| --- | --- | --- | --- |

Observation

| Glassware ID No. | Weight of empty glassware in gm=(A) | (A) + distilled water + (B) | Actual weight of distilled water = (B-A) | Volume in ml=(B-A) C | Average Volume = (V) | Observed tolerance = (V) – specified capacity | Tolerance Limit | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

ENCLOSURE 2

Glassware calibration record

Doc. No.; BML/QDG/QA/039/F01

| Date | Testing Temperature | Type of Glass | Reading (ml/g) | Deviation | Specification | Remark | Done by | Checked by |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

--------------------------

QA Manager