- OBJECTIVE:

To describe the procedure to ensures that the internal temperature and pressure readings are accurate and that the unit consistently achieves the necessary conditions for effective sterilization using biological and chemical indicator

- SCOPE:

To provide written procedure for calibration of Bucket Autoclave in the microbiological laboratory at both infusion and syrup plant in Biomedical Ltd.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Microbiologist: responsible for the operation of the instrument.

- Section in-charge: responsible for reviewing the procedure

- QC Head: responsible for implementing the procedure.

- QAM: overall compliance

- ABBREVIATION AND DEFINITIONS

QC- Quality Control

QAM- Quality Assurance Manager

- PROCEDURES

CALIBRATION:

- Ensure that the equipment is placed in appropriate place

- Before starting the cleaning procedure, disconnect the power supply to the Autoclave.

- Check the bolt of the Autoclave whether is intact

- Check the Air cock of the Autoclave

- Check the pressure gauge.

- Fill the bucket autoclave with water, make sure it cover the element

- Place chemical indicators on all items to be sterilized. Place BIs in "worst-case" or cold spots within a standard, properly loaded basket (e.g., top, middle, and bottom).

- Operate the autoclave through its normal sterilization cycle (e.g., 121°C at 15 PSI for 15-20 minutes).

- After the cycle and cooling, check the chemical indicators. The tape should have visible dark lines, indicating it reached the required temperature.

- Aseptically remove the BI strips and place each into a separate tube of SCDM.

- Include a positive control (an un-autoclaved strip in medium) and a negative control (sterile medium without a strip).

- Incubate all tubes at 55°C-60°C for up to 7 days.

- The positive control shows growth (turbidity), but all test vials and the negative controls show no growth.

- Growth in the test vials indicates a failure in the sterilization process, and the autoclave needs maintenance

- Frequency: This process should be performed every 3 or after major maintenance to ensure consistent results.

- REFERENCE AND ENCLOSURE

Sop for cleaning and operation of Bucket Autoclave – BML/MCB/

7.0 TRAINING INVOLVEMENT

| Quality Control Manager | QC Microbiologist |
| --- | --- |
| QA Doc. Unit staff | Lab. Assistant |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |