- OBJECTIVE:

To describe the procedure of Calibration of Conductivity Meter.

- SCOPE:

To provide written procedure for calibration of Conductivity Meter

3.0 RESPONSIBILITY/ACCOUNTABILITY:

Water Treatment Officer, QA Officer, Chemist, Quality Assurance Manager.

- ABBREVIATIONS AND DEFINITIONS:

N/A

- PROCEDURE:

# CLEANING PROCEDURE

- Switch OFF the mains power unto the Instrument.

- Use dry lint cloth to clean the Conductivity meter daily.

- Clean the electrode after each use with distilled water.

- Store the electrode in cooled MCDP water after rinsing thoroughly.

CALIBRATION PROCEDURE

- Prepare standard solution of Potassium Chloride by dissolving 0.0149gm in 1000ml of distilled water.

- Rinse the electrode several times with WFI/MCDP water.

- Rinse the electrode at least twice with the Potassium Chloride standard solution to be used for the calibration.

- Dip the rinsed electrode in the standard solution and measure the resistivity of the solution at 25.0 + 20C turning the panel screw to 26.6 micro siemens.

- Rinse the electrode thoroughly with distilled water.

REFERENCE/ENCLOSURE:

In-house

- TRAINING INVOLVEMENT

| Chief Engineer | Engineer-Incharge, Operators/Assistants. |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | WATER TREATMENT UNIT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |