- Objective:

To describe the process of Cleaning and operation of Digital pH Meter

- Scope:

This SOP covers the Cleaning and operation of the Digital pH Meter

3.0 Responsibility/accountability:

Quality Control Chemist, Quality Control Microbiologist, Quality Assurance Manager and the Superintendent Pharmacist/ Factory Manager

4.0 Abbreviations and Definitions

PPE - Personnel Protective Equipment

- Procedure

GENERAL:

The Digital microprocessor-based pH meter is supplied complete with a glass body combination electrode, a temperature probe, an electrode holder and 20ml each of pH 4.01 and 7.01 buffer solutions, an electrolyte solution, a 12DC power adapter and an instructional manual.

The equipment is assembled according to the instructional manual and connects the system to a power source.

Cleaning

Daily Cleaning

- Rinse the electrode with distilled water to remove any residual sample.

- Wipe the external body of the electrode gently with lint-free tissue or soft cloth.

- If soiled, clean the electrode with electrode cleaning solution (as recommended by the manufacturer).

- Rinse thoroughly with distilled water after using the cleaning solution.

- Do not wipe the glass membrane; allow it to air-dry if needed.

- Inspect the electrode for any visible deposits or damage.

- Remove and clean the temperature probe with a soft cloth dampened with distilled water.

- Check the electrolyte level inside the electrode and refill or replace if necessary.

- Calibrate the meter using standard buffer solutions after cleaning.

- Record the activity in the Digital pH meter Cleaning Logbook (BML-QDF-CL-009-F02)

#### Sample Measurement Procedure

- Ensure calibration has been performed and the electrode is functional.

- Place the test solution into two beakers: one for measurement and one for rinsing.

- Immerse the electrode and temperature probe at least 4 cm into the sample solution.

- Allow the system to auto-compensate for temperature and wait for the reading to stabilize.

- Record the stabilized pH value from the LCD screen.

- After measurement, rinse the probes with distilled water or cleaning solution.

- Insert the electrode into its protective cap containing storage solution and store properly.

- Record the activity done in the Digital pH meter Usage Logbook (BML-QDF-CL-009-FO2)

Trouble Shooting

- Noise: This causes fluctuating readings. Refer to manual for cleaning procedure or replace electrolyte within glass chamber of electrode

- Dry membrane/junction: soak in storage solution for 1 hour.

- Drifting:  Soak in warm filling solution (50-60OC) for 1 hour and then rinse with distilled water. Refill electrode chamber with filling solution

- Error messages: Errors messages for wrong buffer might require a buffer change, persistent error messages even after this and cleaning procedures mean the electrode is faulty and needs replacement.

Required Materials:

-	Personnel Protective Equipment’s (PPE)

- Distilled water

- Reagent bottles

6.0 REFERENCE/ENCLOSURE

IN-HOUSE

ENCLOSURE 1- DIGITAL pH METER CLEANING LOGBOOK – BML-QDF-CL009--F01

ENCLOSURE 2- DIGITAL pH METER USAGE LOGBOOK – BML-QDF-CL-009--F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QC Analyst | Laboratory Assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |