- Objective:

This SOP describes the procedure for the Cleaning, operation and calibration of the polarimeter.

- Scope:

This procedure covers the Cleaning, operation and calibration of the polarimeters used to determine the optical rotation of samples.

- Responsibility/accountability:

Laboratory Assistant, Quality Control Chemist, Quality Assurance Officer, Quality Assurance Manager.

4.0 Abbreviations and Definitions

-cGMP – Current Manufacturing practice

-SOP – Standard operating Procedure

-PPE – Personal Protective Equipment

- Procedure

#### 5.1 General Precautions

- Handle all components of the polarimeter with care.

- Use only clean and dry polarimeter tubes.

- Ensure samples are at room temperature before analysis.

- Remove all air bubbles from the sample tube prior to measurement.

- Always switch off the instrument after use.

#### 5.2 Cleaning Procedure

- Turn off and unplug the equipment from power supply.

- Wipe the external surfaces using a clean, dry lint-free cloth.

- Clean the polarimeter tube with distilled water and dry it properly.

- Clean the sample chamber gently using soft cloth.

- Document the activity in the Polarimeter Cleaning Logbook (BML-QDF-CL-017-F02).

#### 5.3 Operation Procedure

##### For Analogue Polarimeter

- Connect to power source and switch on the polarimeter.

- Allow the instrument to warm up for 15 minutes.

- Rinse and fill the polarimeter tube with WFI.

- Insert the tube into the sample holder and zero the instrument.

- Fill the tube with the sample solution and insert it into the instrument.

- Read the optical rotation and subtract blank value if necessary.

- Record readings in the Polarimeter Usage Logbook.

##### For Digital Polarimeter

- Switch on the instrument and allow 15 minutes warm-up.

- Rinse and fill the polarimeter tube with blank solution.

- Place tube in the sample chamber and press the START button.

- Repeat the blank measurement 3 times and calculate the average.

- Record this value.

- Fill the tube with the sample solution, press START to take reading.

- Repeat the process 3 times and determine the average.

- Subtract the blank from the sample reading.

- Document results accordingly in the Polarimeter Usage Logbook (BML-QDF-CL-017-F02).

#### 5.4 Calibration Procedure

- Dry sucrose in an oven at 105°C for 4 hours.

- Prepare sucrose solutions at 5%, 10%, and 26% w/v.

- Rinse and fill the polarimeter tube with distilled water to take a blank reading.

- Clean and fill the tube with sucrose solutions in increasing order.

- Take three readings for each solution and calculate average.

- Determine the specific optical rotation using standard formulas.

Materials, Equipment, Apparatus and Reagents

- PPE, sucrose solution, volumetric flask, pipette, wash bottle

- Crucible

- Oven

6.0 REFERENCE/ENCLOSURE

Equipment manual

ENCLOSURE 1- POLARIMETER CLEANING LOGBOOK- BML-QDF-CL-017-F01

ENCLOSURE 2- POLARIMETER USAGE LOGBOOK- BML-QDF-CL-017-F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QC Analyst | Laboratory Assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |