- OBJECTIVE

To provide written guidelines for cleaning, operation and calibration of Microprocessor Touch Screen Digital Polarimeter

- SCOPE:

The operation and calibration of Microprocessor Touch Screen Digital Polarimeter Instrument to determine the specific optical rotation of given test sample.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- The Chemical Analyst shall strictly follow this procedure to achieve safe and efficient use of equipment.

- The user shall ensure the equipment is kept clean and reports any malfunctioning.

- The Quality Control Head ensures scheduled maintenance is carried out as at when due.

- The Quality Assurance Manager ensures full compliance in using the equipment.

4.0 DEFINITIONS AND ABBREVIATION:

- LGB – Logbook

5.0 PROCEDURE:

Materials, Equipment, Apparatus and Reagents

- Sucrose solution

- Volumetric flask

- Pipette

- Washing bottle

- Tissue paper

- Measuring tube

Precautions

- Handle the instrument and the Microprocessor Touch Screen Polarimeter tube carefully.

- Do not keep Microprocessor Touch Screen Polarimeter tube without cleaning after use; it must be clean after use.

- Before putting the equipment on, ensure it is connected to UPS ( uninterrupted power supply)

- Keep switch off the instrument when not in use.

- Specific optical rotation to be measure sample should be clear.

- Avoid air bubbles in the sample tube and if any air bubble observed , make the air bubbles float on the protruded tube neck

- The screw-nut of the test tube should not be screwed down too tightly, in order for the readings not to be affected.

- Always Keep the Polarimeter tube in its box after use.

Calibration

- Dry the sucrose at 105oc for 4 hours and prepare 10.0% w/v, 20.0% w/v, 30.0% w/v, 40.0% w/v and 50.0% w/v.

- Fill the Microprocessor Touch Screen Polarimeter tube with WFI water at specific temperature( 250c)

- Ensure the optical rotation read Zero before use or press zero button on display before measurement.

- Remove the tube and pour the water sample out, rinse the Microprocessor Touch Screen Polarimeter tube with prepared 10.0%w/v sucrose solution and then fill it with 10.0%w/v sucrose solution.

- Allow to set and take the reading.

- Repeat above procedure for 20.0% w/v, 30.0% w/v, 40.0% w/v and 50.0% w/v sucrose solution and record the data in format calibration of Microprocessor touch screen Microprocessor Touch Screen Polarimeter and affixed calibration label.

- In-house Calibration frequency: quarterly

- External calibration : twice a year

## Operation

- Switch on the mains and let the equipment warm up 10 minutes.

- Ensure the Polarimeter tube is clean before use

- Touch the screen and remove the Polarimeter tube from the Microprocessor Touch Screen Polarimeter

- Fill the Polarimeter tube with WFI water at 250c and place inside the Microprocessor Touch Screen Polarimeter chamber, then press “ZERO” to Blank the equipment.

- Pour the wfi water away, then rinse the Polarimeter tube with the test sample at specific temperature.

- Fill the Polarimeter tube with the test sample at specific temperature, the sample must be clear and free from air bubble.

- Clean the sides of the tube with tissue paper. If any air bubble comes then it must be trapped in the center of Microprocessor Touch Screen Polarimeter tube.

- Keep the Polarimeter tube in its place in the Microprocessor Touch Screen Polarimeter chamber.

- Press measurement reading, take the reading

- After completion of reading, take out the Polarimeter tube from its chamber and clean. Keep it in its box.

- Then switch off the instrument.

- Document all analysis activity in the instrument usage logbook and analyst workbook.

Safety, Health and Environment Precautions

- Wear appropriate PPE

6.0 REFERENCE AND ENCLOSURE

- Equipment Usage Logbook – BML-QDF-CL-027-F01

- Equipment Calibration Form – BML-QDF-CL-027-F02

- Instrument Manual.

- TRAINING AND INVOLMENT

| QUALITY MANAGEMENT | OTHERS |
| --- | --- |
| Quality Assurance Manager | Chemical analyst, IPQA |
| Head of the laboratory | Microbiologist |

8.0 DISTRIBUTION AND CONTROL COPY

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | NIL | First issue |