- OBJECTIVE:

To provide a procedure for calibration of Incubators to ensure that the temperature maintained inside the incubator is accurate and uniform for microbiological incubation

- SCOPE:

This procedure applies to all incubators used in the microbiology laboratory for incubation of culture media

3.0 RESPONSIBILITY/ACCOUNTABILITY:

Laboratory Attendant, Quality Control Microbiologist, Quality control Head, Quality Assurance Manager, Superintendent Pharmacist.

4.0 ABBREVIATIONS AND DEFINITIONS:

N/A

- Procedure

- Ensure the incubator is clean and empty

- Verify that the reference thermometer is calibrated and within validity period

- Switch ON the Incubator

- Set the incubator temperature according to its operating range i.e 30-35°C for bacteria and 20-25°C for fungi

- Place calibrated thermometers at different positions inside the incubator

- Close the incubator door properly

- Record the displayed incubator temperature

- Record the actual temperature from the reference

- Take readings at 30-minutes interval for 2 hours

- Record all readings in the calibration logbook in the format below (BML.QA.MCB.ECR.001B)

- The incubator temperature must be within ±2°C of the set temperature

- The calibration shall be carried out quarterly

| Set Temperature | Time | Thermometer | Thermometer Reading (°C) | Incubator Reading (°C) | Tolerance (°C) |
| --- | --- | --- | --- | --- | --- |

Safety Precaution

- Do not open the incubator frequently during calibration

- Use only calibrated thermometers

Materials, Equipment, Apparatus and Reagents

- Calibrated reference thermometer

- Incubator

- Calibration Logbook

- Stopwatch

- REFERENCE/ENCLOSURE:

REFERENCE: EQUIPMENT MANUAL

ENCLOSURE 1- INCUBATOR CALIBRATION LOGBOOK - (BML.QA.MCB.ECR.001B)

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY:

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |