- OBJECTIVE:

This SOP outlines the procedures for monitoring and controlling the microbiological bioburden in pharmaceutical production areas to ensure a clean and safe manufacturing environment.

- SCOPE:

- This procedure applies to all production areas within the pharmaceutical manufacturing facility.

- RESPONSIBILITY/ACCOUNTABILITY:

- Microbiologist Perform the microbiological bioburden monitoring as per the SOP.

- Quality Control Head ensures compliance with the SOP and review test results.

- ABBREVIATIONS AND DEFINITIONS:

- CFU: Colony Forming Unit

- Bioburden: The number of microorganisms in the air of a production area.

- PROCEDURE

- Sterilize all glass wares needed for the environmental bioburden in the autoclave according to SOP BML-QDS-MC-007

- The work bench and the Laminar Flow Cabinet should be properly disinfected with 70% IPA / as per SOP BML-QDS-MC-008

- Prepare the TSA and SDA media required for plate exposure as per SOP BML-QDG-MC-001

- Pour about 15ml to 20ml of sterilized media cooled to 450c aseptically into different sterile Petri-dishes under Laminar Air Flow and allowed to solidify.

- After solidification, pre incubate all the media containing petri dishes for 24 hours in inverted position for detecting any contamination.

- After pre-incubation, examine all the petri plates, expose the plates without any contamination and discard any contaminated plate according to SOP BML-QDG-MC-008.

- Plates should be labeled accordingly

- Sanitize the sampling box with 70% Isopropyl Alcohol

- Take all pre-incubated plates containing media in the sampling box to production areas for exposure.

- Expose the petri plates at each location by removing the lid of the petri plate and placing the plate containing the media at a point.

- Expose the petri plates for 4 hours

- After exposure for 4 hours, close the petri plate with lid, collect all the petri plates in the sampling box and carry all exposed plates to laboratory for incubation.

- Incubate all exposed Sabouraud Dextrose Agar plates at 20-25oc for 72 hours and all Tryptic Soya Agar plates at 30-35oc for 48 hours in inverted position.

Quality Control:

- Include positive and negative controls in each batch of tests to ensure the accuracy and reliability of the results.

Results Enumeration

- After incubation, count the number of colonies (CFU) on each plate.

- Record the results, noting the location and type of media.

- Compare the microbial counts (CFU) with acceptance criteria specified in relevant standards or internal guidelines.

- Investigate any counts that exceed the acceptable limits and take appropriate corrective actions.

- Record all observations, test results, and any deviations in the appropriate logbooks

- Ensure all documentation is signed and dated by the person performing the test and reviewed by the Quality Assurance Manager.

Reporting

- This should be done if the Microbial level of the area is confirmed high or above the stated limit.

WHO limit ------NMT 200 cfu

Alert limit ------- NMT 50 cfu

Action limit ------NMT 100 cfu

Corrective Action

- Fumigation exercise should be done immediately.

Materials and Equipment:

- Sterile Settle plates (Petri dishes with agar)

- Sabouraud Dextrose Agar and Tryptic Soy Agar

- Incubators

- Autoclave

- Laminar Flow Cabinet

- Sampling box

- 70% isopropyl alcohol

- Personnel protective equipment PPE

Safety:

- Follow all laboratory safety protocols, including the use of personal protective equipment (PPE).

- Dispose of all biological waste in accordance with local regulations and company policies.

- REFERENCE/ENCLOSURE

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |