- OBJECTIVE

To outline procedure for cleaning operation, calibration, verification and cleaning of 150kg maximum capacity platform scale.

- SCOPE:

This procedure applies to operation, calibration, verification and cleaning of the 150kg scale stationed in the raw material store, for its processes, and other necessary activities that have the potential to affect cGMP.

- RESPONSIBILITIES/ACCOUNTABILITIES:

Store and IPQA personnel are responsible for ensuring that this SOP is followed for the operation, calibration, verification and cleaning of the platform scale.

- DEFINITIONS AND ABBREVIATION:

  - cGMP: Current Good Manufacturing Practices

  - QA: Quality Assurance

  - IPQA: In-Process Quality Assurance

- PROCEDURE:

CLEANING

- Cleaning should be done after every use.

- Switch off the balance.

- Disconnect balance from mains supply.

- First remove powder and dust, then sticky substances.

- For powder and dust use a dry napkin. Never blow; this could transport dirt or spilled sample materials inside the balance.

- For removal of sticky substances, use a damp lint-free cloth and mild solvent such as 70% ethanol; avoid abrasive materials.

- Do not spray or pour liquids directly on the scale platform.

- Remove parts for cleaning where possible (e.g. weighing platform). Only remove parts that can be removed without tools.

- If possible, do not disconnect peripheral devices such as printer unless cleaning is not possible without doing so.

Clean the balance in its working location; do not tilt, move or carry it around unnecessarily. Incorrect handling may cause expensive damage.

OPERATION

    - Ensure that the platform pan is clean & free from dust.

- Connect plug to mains and switch ON the balance by pressing the holding the Power button for 2 seconds.

- Press the →0← button to zero the weight display. The scale must be stable to accept zero operation.

- To tare the scale, place a container on the platform and press →T← button.

- To weigh, place the item to be weighed on the container.

- As a material is added only the weight of the material will be displayed.

- After weighing, remove all materials from top of the platform and press the →0← button to zero the scale.

- Clean the scale immediately after use.

- Sign the weighing scale usage logbook LGB/WHF/WR/001/01

| Date | Used for | Start time | End time | Total time used | Status | Operator Sig. | Checked by QA | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

CALIBRATION

- Access the Calibration menu by pressing the little black button on the motherboard inside the panel containment.

- Empty the scale and press the ENTER button to start calibration. “E SCL” (empty scale) appears on the display.

- Press ENTER again. The display will count down 10…0 as the terminal captures zero.

- ”Full Ld” appears on the display.

- Press the ENTER button. The full capacity value appears, blinking.

- Use the →T← or →O← buttons to move down or up respectively to change the calibration weight value.

- When the value has been set, put the corresponding full capacity weight on the scale and press the ENTER button.

- The display will count down 10…0. If the calibration was successful, “done” will display for 2 seconds.

- After calibration, the calibration should be verified using standard weights 5kg, 20kg and 100kg

- Repeat calibration if tolerance or deviation is outside ±0.1.

- Record displayed values of the weights in the verification and calibration logbook LGB/WHF/WR/001/02

| Date | Standard weight used (kg) | Readings | Deviation | Specification + 0.1Kg | Remark | Done by Operator | Checked by QA |
| --- | --- | --- | --- | --- | --- | --- | --- |

VERIFICATION

- Verification shall be done daily.

- Before verifying the scale, clean the weighing platform.

- Ensure the level indicator is at the centre of the circular marking.

- Switch ‘ON’ the main power supply.

- Wait till the zero reading is displayed on the screen.

- Allow to stabilize for 15 -20 seconds before placing standard weights on pan.

- Place 5kg, 20kg and 100kg standard weights consecutively on the platform.

- Allow to stabilize before taking reading.

- Affix Verification label on the scale.

6. REFERENCE AND ENCLOSURE:

- BML/WHF/WR/001/F01 - cleaning log

- BML/WHF/WR/001/F01 – calibration log

- BML/WHF/WR/001/F01 – Usage log

7.0 TRAINING INVOLVEMENT

| All store officer | Store unit head |
| --- | --- |
| All QA/QC officer | All IPQA officers |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | STORE |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | - First issue |
| 01 | 00 | - To review |
| 05 | 04 | - Due for review |
| 06 | 05 | - Due for review - Inclusion of intermediate weight |