- OBJECTIVE:

To provide written guidelines for Cleaning and operation of Water Bath

- SCOPE:

This SOP covers the process of Cleaning and operation of Water Bath

- RESPONSIBILITY/ACCOUNTABILITY:

- The user must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS

- Nil

- PROCEDURE

Cleaning

- Switch ‘OFF’ the mains

- Remove the lid from the water bath

- Wash the lid properly with purified water

- Drain the water from the water bath by opening the valve at the back, wash the water bath with DM water and mop with a clean dry cotton cloth

- Fill DM water up to 3/4th height of the bath.

- Put the lid of the water bath

Note: in between when the water becomes dirty follow the same from the start

Operation

- Ensure that the instrument is clean and calibration of temperature indicator is within due date of calibrate

- Switch “ON” the ring both by pressing “ON/OFF” switch.

- The digital temperature controller and indicator will indicate the actual temperature of water

- Set the desired temperature by pressing SET switch and adjusting the temperature.

- If the indicated temperature is less than desired, press the up or down button.

- Upon completion, fill the water bath user logbook provided.

Water bath User Logbook LGB/QCS/CL/017/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

PRECAUTIONS

- Do not operate without water inside the instrument

- Switch OFF when the instrument if not in use

Materials, Equipment, Apparatus and Reagents

- PPE’s

- Demineralised Water

- REFERENCE/ENCLOSURE

- Bright Medical water bath instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Chemical Analyst |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |