- OBJECTIVE:

To provide written guidelines for Cleaning and Operation of Refractometer

- SCOPE:

This SOP covers the process of Cleaning and Operation of Refractometer

- RESPONSIBILITY/ACCOUNTABILITY:

- The Chemical analyst shall safely and properly handle the equipment.

- The user reports any deviation or equipment malfunction.

- The equipment user must ensure that the equipment is always kept balanced and clean.

- The Quality Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS

- Nil

- PROCEDURE

CLEANING PROCEDURE:

- Apart from daily routine cleaning, the equipment must be cleaned once a week.

- Open and clean the prism.

- Clean the outer part of the apparatus.

- Clean the Apparatus from its original position clean the base and then replace

Testing work:

To test transparent and translucent liquid:

- Drop the tested liquid onto the surface of the prism with a dropping tube then cover the incident prism and lock with the knob

- With the demand that the liquid layer should be homogeneous, spread full of the field view, bubble less.

- Open the shield and close the reflecting mirror, then adjust the ocular dipoter so that the cross - line image sharply, then turn the knob to look for the position of the bright and dark boundary line in the ocular field of view,

- Turn the fine-focus the knob to make the boundary line lie in the center of the cross - lines, and eventually turn the condenser properly.

- At this moment the value indicated below the ocular field of view is right the refractive index of the liquid under test.

To test transparent solids:

- The object to be tested needs a smooth polished surface.

- Open the incident prism and add one or two drops of naphthalene bromide onto the smooth plane of refracting prism, them clean the polished surface of the tested object so that it can contact better, when the work looking for the boundary line in the eyepiece can be conducted.

- The procedure for aiming and reading accords with the above - mentioned.

To test translucent solid:

- One surface of the translucent solid should be a polished plane, upon which naphthalene bromide should be dropped, then stick it on the refracting prism.

- Open the reflecting mirror and adjust the angle and use it as a surface of incidence for measuring, the method for operation being the same as the above - mentioned.

To test the sugar concentration of sugary liquid:

- The operation is similar with testing the refractive index of liquid, its reading can be taken from the indicated value of the field of view,

- The readout taken then is right the percentage of the sugary liquid concentration.

To determine mean dispersion value:

- Turn the knob until the color disappears in relation to the bright and dark boundary line in the field of view, when the indicated value Z in the scale knob for dispersion graduation and its refractive index nD should be taken down.

- According to nD, find the corresponding value of A and B from the dispersion table.

- If nD lies in between two figures, it can be found by interpolation.

- Then find the corresponding value of O' from the dispersion table according to the value Z.

- When Z>30, the negative sigh; when Z>30, take positive sigh.

- The mean dispersion value can be given by substituting the found A,B and O' into the dispersion equation

- When refractive indices are required to measure at various temperatures, insert the thermometer into its socket  and connect it,

- Adjust to the temperature required, and the measurement can be conducted after the temperature is stabilized for the minutes.

- Record the usage activities in the user logbook for the Refractometer appropriately.

Refractometer User Logbook LGB/QCS/CL/024/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

- REFERENCE/ENCLOSURE

- Refractometer instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Chemical Analyst |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |