- OBJECTIVE:

This SOP describes the procedure to be followed in operating high performance liquid chromatography(HPLC).

- SCOPE:

This procedure is applicable to the use of high-performance liquid chromatography(HPLC).

- RESPONSIBILITY/ACCOUNTABILITY:

  - The Chemical Analyst shall strictly follow this procedure to achieve safe and efficient use of equipment

  - The user shall ensure the equipment is kept clean and reports any malfunctioning.

  - The quality control manager ensures scheduled maintenance is carried out

4.0 ABBREVIATIONS AND DEFINITIONS

  - HPLC: an acronym which stands for HIGH PERFORMANCE LIQUID CHROMATOGRAPHY (HPLC) which refers to high resolution, sensitivity and efficiency- a method of chemical separation of compounds.

- PROCEDURE

Operation

- Put on the unit from the switch at the back, ensure each individual module’s power switch            illuminates in red.

- Put on each individual each individual module from the switch illuminated in red ensure the LCD comes on and there’s no error display on each module.

- Rinse the pump using water rinsing should be done before and after any analysis especially if you are working with buffers.

- Prime the pump by setting each flow line (A, B, C, D) individually to 100% from the keypad on the pump and press “enter”, turn the prime valve two full turns anticlockwise place a syringe in the valve hole and draw using the plunger (priming should be done in situation where the flow lines are completely dry of solvents) using HPLC grade water.

- Install the column; do not connect the outlet, place end in a beaker.

- Put the mobile phase for the current analysis in the reservoir water should always be in Position A

- Set the mobile phase combination for the analysis, set a flow rate of 0.5ml/min and activate the pump, watch the pump outlet and ensure fluid is dropping, after 1-3 min. stop pump connect outlet to detector and activate pump again.

- Open software

- Create a new method or open a pre-existing method and verify parameters for analysis (Flow rate, detector wavelength, oven temperature) and “send method”

- Access device monitor from software turn on detector lamp

- Wait for 10-15 minutes for detector to warm up and stabilize

- Rinse syringe in methanol and HPLC grade water

- Draw sample for injection

- Set manual injector to “load position”

- Inject sample and turn knob to “inject position”

- Access Data Acquisition from software and monitor the analysis.

- After analysis stop the pump, if working with buffers rinse pump head with water

- If the previous analysis was done with a buffer transfer all flow line into water and reduce flow rate to 1ml/min and run for 15-20 minutes.

- After flushing with water transfer flow lines to reagent with a 70% Methanol and 30%water and flush for another 10-15 minutes.

- Upon completion of the reading, fill the HPLC user logbook provided.

HPLC User Logbook LGB/QCS/CL/026/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

- If not using the instrument for up to 2 weeks remove column, seal ends and store in supplied box.

- Transfer the flow lines to the individual reservoir bottles and store.

- Turn off the unit.

Cleaning procedure

For system cleaning

- After completion of analysis, Clean the HPLC as follows

- Transfer the used glassware for washing and keep aside test tubes stands lying near the HPLC column. Clean the surrounding area with tissue paper

- Clean the syringe with repeated rinses of distilled water

- Clean the outer surface of injector by flushing the distilled water with syringe forcefully near the injector hole and joints of loops. Clean with tissue paper

- Clean the inner surface of injector and loops by injecting distilled water into the loop syringe three to four times

- Clean the plunger seal by injecting distilled water three to four times with syringe

- Remove the column and replace with joints for the cleaning of tubing and flush with 6N nitric acid, distilled water and then with methanol

- Remove the air suction pipe by purging the line and flushing with distilled water for some time

- Clean the suction filter by sonicating it into nitric acid for 10- 15minutes and the wash with distilled water

- Replace the joints with column and wash the column first with distilled water then with 30% methanol

- Switch OFF the pump, detector and integrator

- Clean the outer surface of the instrument with dry clean cloth

- Remove any dirt or spot with Isopropyl alcohol

- Dry with tissue paper

Materials, Equipment, Apparatus and Reagents

- Thermometer

- Cotton cloth

- Sonicator

- Nitric Acid

- Isopropyl alcohol

- Methanol

- Distilled water

- Tissue Paper

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

    - HPLC instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Chemical Analyst |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |