- OBJECTIVE:

To provide written guidelines for cleaning and Operation of Electronic Weighing Balance in syrup microbiology laboratory.

- SCOPE:

This SOP covers the process of cleaning and Operation of Electronic weighing Balance used in Microbiology Laboratory.

3.0 RESPONSIBILITY/ACCOUNTABILITY:

- The microbiologist shall carry out calibration procedures using standard weights.

- The Quality Control Head reviews the results as against the specifications.

- Users shall strictly follow this procedure to achieve safe and efficient use of equipment

- The user shall ensure the equipment is kept clean and reports any malfunctioning.

- The Quality Control Head ensures scheduled maintenance is carried out.

4.0 ABBREVIATIONS AND DEFINITIONS

- Calibration: to mark of gauge an instrument with a standard scale of readings

- IPA -   Isopropyl alcohol

- PPE -   Personal protective equipment

- PROCEDURE

Safety Precaution

- Ensure the Weighing balance is switched OFF before cleaning.

- Check the main power source and ensure it is OFF.

- Do not move or change the position of the weighing balance while cleaning

Cleaning Procedure

- Open the side glass and clean the top load and inside cubicle with a dry lint-free cloth daily so that there will be no dust particles on the weighing balance.

- Clean the top glass and side glass with dry cloth.

- Re-clean the inside cubicle and top load with 70% IPA moist cloth and wipe dry after cleaning.

Operation Procedure

- Ensure the cleanliness of the instrument.

- Ensure for the level of weighing balance, air bubble shall be position in the center.

- Switch on the main power source and the UPS attached to the balance.

- Switch on the balance from the adaptor power unit and press the T button on the balance panel to power it on.

- Allow the weighing balance to warm up for 15 minutes.

- Calibrate the weighing balance by following the procedure in the calibration SOP

BML-QDG-QA-032.

- When weighing do not weigh anything more than 500g on the balance.

- Place materials to be weighed on foil paper with a clean spatula/spoon up to the desired target weight.

- To tare the weighing balance, press the tare button.

- After using the weighing balance ensure it’s clean and not soiled by any media or materials.

- Log in all activities in the user’s logbook (LGB/QDS/MC/005/01)

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

Materials, Equipment, Apparatus and Reagents

- PPE

- Foil paper

- Spatula

- Cotton wool

- 70% IPA

Safety, Health and Environment Precautions

- Be prompt to clean any spillage

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

- Analytical weighing balance instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Microbiologist, Laboratory assistance |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |