- OBJECTIVE:

To provide written guidelines for cleaning and Operation of Lamina Air Flow to maintain a sterile environment.

- SCOPE:

This SOP covers the process of cleaning and Operation of Lamina Air Flow

- RESPONSIBILITY/ACCOUNTABILITY:

Laboratory Attendant, Quality Control Microbiologist, Quality Assurance Officer, will ensure adherence to this SOP for the operation and cleaning of the Laminar Air flow Chamber.

4.0 ABBREVIATIONS AND DEFINITIONS

UV – Ultra-violet

- PROCEDURE

Safety Precaution

- Ensure the Lamina Air Flow is switched OFF before cleaning.

- Check the main power source and ensure it is OFF

- Do not MOVE or change the position of the UV light and the light bulb

Cleaning Procedure

    - Clean the work surface and wall with appropriate disinfectant e.g. 70% ethanol and Oxonian solution.

    - Use damp cloth to clean the exterior surface of the lamina airflow and work bench.

    - Use methyl ethyl ketone to remove stubborn stains and spots on the stainless-steel		surface [ if any]

    - Wash the steel surface immediately with 1% w/v detergent solution

    - Wipe using soft cloth soaked in water.

    - Do not use chlorine containing disinfectant.

Operation Procedure

- Connect to the main electricity supply.

- Put on the ultraviolet ray fluorescent by pressing down the UV knob.

- For disinfection of the chamber, 30mins to 1hour of exposure to UV light is adequate.

- Switch off the UV lamp after the disinfection time.

- Switch on the normal light and airflow. The airflow supplies sterile air.

- It is advisable not to put on the UV light while working to prevent possible photo reactivation.

- Carry on the operation in the hood as fast as possible.

- Switch off the machine after the operation and disconnect from the main electrical supply.

- Clean the machine after the operation.

- Record the usage activities in the user logbook LGB/QDS/MC/008/01 for the Laminar Air Flow appropriately.

Laminar Air Flow User Logbook LGB/QDS/MC/008/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

Materials, Equipment, Apparatus and Reagents

- PPE

- Cotton wool

- 70% IPA

- REFERENCE/ENCLOSURE

- NIL

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |