- Objective:

To provide written instruction on the cleaning and operation of the wet and dry hygrometer

- Scope:

To provide written guidelines for the cleaning and operation of the wet and dry hygrometer

3.0 Responsibility/accountability:

Q.A. Officer, QC chemist, Microbiologist, Q.A. Manager

4.0 Abbreviations and Definitions

N/A

- Procedure

EQUIPMENT CLEANING

  - Check the Hygrometer for any leakage at the wet bulb water holder.

  - Clean any dirt or water on the Hygrometer with a clean dry cloth

  - Ensure that the capillary tubes are well cleaned daily.

  - Clean the hold slate of the Hygrometer daily.

CALIBRATION

  - Take a calibrated thermometer and wrap round the bulb small quantity of cotton wool

  - Immersed the cotton wool in cool clean water

  - Wave the thermometer gently round the area where your hygrometer is placed for 10minutes and take your readings

  - Remove the wet cotton wool and wipe dry the bulb of your thermometer

  - Wave the dry thermometer gently round the area for at least 10 minutes and take your reading

  - Calculate your depression by subtracting the wet reading from the dry reading of your thermometer

  - Check the reading of the Hygrometer placed in the area, any deviation of more than, ±2.0OC for temperature and ±5%for humidity is not acceptable. If any such deviation is observed check if the Hygrometer wet bulb water cup holder is clean

  - Check if the dry bulb is cleaned and not broken

  - Report any such deviation to the Q.A. Manager

OPERATION

  - Fill the water holding cup Hygrometer at the wet bulb end daily.

  - Ensure that the dry bulb is free of any dirt attached to it

  - Take the reading on the dry bulb first, and then read the wet bulb.

  - Calculate your depression of the wet bulb by subtracting the reading of the dry bulb from the  reading of the wet bulb

  - Check the chart attached to the slate and read on the dry bulb the temperature, check for the corresponding reading of the depression for your humidity reading.

  - Record your readings in the chart for daily recording of the temperature and humidity.

  - Report any abnormal temperature and humidity immediately to the Q.A. Manger for necessary action.

- REFERENCE/ENCLOSURE

In House Method

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 02 | ### 01 | ### July 2022 | To update |
| ### 03 | ### 02 | ### July 2024 | - Inclusion of enclosures format - Change in sop format generally - To update |