- Objective:

To lay the procedure for the calibration of the conductivity meter.

- Scope:

This SOP covers the procedure for the calibration of the conductivity meter (BML/QAE/SYP/008)

- Responsibility/accountability:

Laboratory Assistants, QA/QC analyst, the Quality Assurance Manager and the Superintendent Pharmacist.

4.0 Abbreviations and Definitions

- GLP: Good Laboratory Practice

- GDP: Good Documentation Practice

- Procedure

- Connect the meter to the powerline, press the button.

- Preheat the meter for 30 minutes.

- Rinse the conductivity probe with distilled water.

- Immerse the probe into the standard solution after having controlled the solution temperature to ambient temperature.

- Read the conductivity value Kc, then calculate the cell constant J. J=K/Kc. (K: standard conductivity of the solution, 1408.3µs/cm at 250C for 0.01M KCl.)

- Setting the cell constant: Press these keys sequentially: SETUP key, UP key, SETUP key, OK key. Input the calculated value J. then press OK, and EXIT.

- Frequency for accuracy: calibrate meter every day, and as when required when drift is noticed.

MATERIALS, EQUIPMENT, APPARATUS AND REAGENTS:

- 0.01M KCl

SAFETY HEALTH AND ENVIRONMENTAL PRECAUTIONS:

- PPE

- REFERENCE/ENCLOSURE

Equipment Manual

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |