- Objective:

To provide written guidelines for Calibration of the Moisture Content Analyzer

- Scope:

This SOP covers the procedure for Calibration of the Moisture Content Analyzer

- Responsibility/accountability:

QA Officer, Chemist, Microbiologist, Quality Assurance Manager and Superintendent Pharmacist /Factory Manager.

4.0 Abbreviations and Definitions

N/A

- Procedure

- Calibration

Weight Calibration

- Built-in balance should be calibrated if the analyzer has been abused, and if this is stipulated    by quality assurance system

- Press “SETUP”, press “BACK”, press “CALIBRATION” and then press “BACK”

- Press “WEIGHT ADJUSTMENT- EXTERNAL”

- place 50g weight, close the upper housing and note the reading

- Follow the screen prompts throughout the process.

Accuracy of Balance

- Place 0.5g weight, close the upper housing and note the reading

- Follow the screen prompts throughout the process.

- Place 1g weight, then repeat the process as stated above and record the observed result.

- then follow the instruction on the display

Temperature Calibration

- Press “SETUP”, press “BACK”, press “CALIBRATION” and then press “BACK”

- Press “TEMPERATURE CALIBRATION”

- Set the temperature 600C

- Remove the pan support, insert the calibrated thermometer from top side to near temperature censor and close the upper housing

- Press the “START” to initiate the temperature calibration process

- After constant temperature achieved, observe the temperature and record your reading

- Follow the screen prompts throughout the process.

- Set the temperature to 1050C and 1500C, then repeat the process as stated above and record the observed result.

Accuracy of Drying

- Open the cover on the Moisture Analyzer.

- Wet the glass fiber pan with a 0.5 to 1 gram of water

- Start the instrument and observe the results

- Repeat the test twice and record the observed result

Frequency: once in three month

Caution

- Be careful when removing the temperature calibration unit from the dryer unit as it can be very hot. Allow it to cool down by opening the cover before removing.

- Do not touch or clean the halogen heater.

- Do not place any combustible materials on, under or next to the instrument.

- The vent over the sample must never be covered, plugged, taped over or tampered with in any other way

- Never leave the instrument unattended!

Materials, Equipment, Apparatus and Reagents

PPE

Safety, Health and Environment Precautions

- Be prompt to clean any chemical spillage

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

IN-HOUSE

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |