- Objective:

To provide written guidelines for calibration of UV Spectrophotometer.

- Scope:

This SOP covers the process of calibration of UV Spectrophotometer.

- Responsibility/accountability:

Laboratory Attendant, Quality Control Chemist, Quality Assurance Manager

Superintendent Pharmacist/ Factory Manager.

4.0 Abbreviations and Definitions

N/A

- Procedure

CALIBRATION:

CALIBRATION OF WAVELENGTH/CONTROL OF ABSORBANCE

- Dry a quantity of potassium dichromate by heating to constant weight at 1300C

- Weigh and transfer accurately a quantity not less than 57.0mg and not more than 63mg (60mg) to 1000mL volumetric flask.

- Dissolve and dilute in sufficient 0.005M H2SO4 to produce 1000mL.

- Measure the absorbance of potassium dichromate solution at wavelengths given below, and then calculate the A1%1cm for each wavelength.

CALIBRATION OF STRAY LIGHT

- Prepare a 1.2g of potassium chloride previously dried at 1050C for two hours into 100ml volumetric flask

- Add water to dissolve and make up to volume with water.

- Measure the absorbance of the solution at 198.0, 199.0, 200.0, 201.0, 202.0nm using water as blank.

- Acceptance criteria: absorbance greater than 2.00

CALIBRATION OF RESOLUTION POWER

- Prepare a 0.02% v/v toluene in hexane.

- Scan through wavelength 260nm to 420nm.

- Acceptance criteria: the ratio of the absorbance at maximum at about 269nm to that at the minimum at about 266nm is not less than 1.5.

Materials, Equipment, Apparatus and Reagents

- PPE

- Toluene

- Hexane

- Potassium chromate

- Potassium chloride

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

British pharmacopoeia

Instruction manual

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |