- Objective:

To provide written guidelines for calibration of Laboratory Oven.

- Scope:

This SOP covers the procedure for Calibration of Laboratory Oven.

- Responsibility/accountability:

Laboratory Assistance, Quality Control Chemist, Quality Control Microbiologist, Quality Assurance Manager, Supt. Pharmacist.

4.0 Abbreviations and Definitions

N/A

- Procedure

Calibration:

- Switch “ON” by pressing “I” on “I or O” switch.

- Set the temperature to 500C

- Put the standard thermometer on the upper shelf of the oven and close the door of the oven

- Repeat the above procedure by putting the thermometer in lower shelf

- Record the observed reading

- The observed temperature of the thermometer in both shelves should be ±2.00C tolerance limit to the set temperature

- Set the temperature to 1050C and 1500C, then repeat the process as stated above and record the observed result.

- Frequency: every three (3) months

Materials, Equipment, Apparatus and Reagents

Thermometer

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

IN-HOUSE

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |