- OBJECTIVE:

This SOP describes the procedure to be followed in Calibration of Incubator

- SCOPE:

To provide written procedure for calibration of Incubator

- RESPONSIBILITIES/ACCOUNTABILITIES:

QA Officer, Microbiologist, QA Manager

- ABBREVIATION AND DEFINITIONS

QC – Quality Control

- PROCEDURES

CALIBRATION PROCEDURE:

- Check the thermostat if it actually set to the required temperatures (i.e. 250C for fungi growth and 370C for Bacteria).

- Check if the metal shelves are cleaned.

- Make sure that the door of the incubator is properly otherwise temperature will leak and accuracy will not be maintained.

- Place at least 4 calibrated thermometer inside the incubator in different places

- Operate the incubator as per the standard operating procedure –BML-QDF-MC-007

- Document your readings from each calibrated thermometer at an interval of 30minutes and record the results up to 2hours

- Record all readings in the calibration logbook in the format below (BML.QA.MCB.ECR.001B)

- The incubator temperature must be within ±2°C of the set temperature

- Report to QC Manager immediately if there is lapses

- Frequency – once in 3 month

| Set Temperature | Time | Thermometer | Thermometer Reading (°C) | Incubator Reading (°C) | Tolerance (°C) |
| --- | --- | --- | --- | --- | --- |

Safety Precaution

- Do not open the incubator frequently during calibration

- Use only calibrated thermometers

Materials, Equipment, Apparatus and Reagents

- Calibrated reference thermometer

- Incubator

- Calibration Logbook

- Stopwatch

- REFERENCE AND ENCLOSURE:

REFERENCE: EQUIPMENT MANUAL

ENCLOSURE 1- INCUBATOR CALIBRATION LOGBOOK - (BML.QA.MCB.ECR.001B)

7.0 TRAINING INVOLVEMENT:

| Quality Control Manager | QC Head |
| --- | --- |
| QC Microbiologist | Lab. Assistant |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | MICROBIOLOGY LABORATORY |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in SOP format. - Inclusion of traceability of enclosures to parent SOP. |