- Objective:

This procedure is to serve as a guide for water sampling and Sterility testing of WFI

- Scope:

- This procedure should be performed daily.

- It is applicable to water for injection (WFI).

3.0 Responsibility/accountability:

- A qualified Microbiologist should carry out this procedure.

- The Quality Assurance Manager should ensure that it is performed.

- Abbreviations and Definitions

TSB – Tryptone Soya Broth

FTM– Fluid Thioglycollate Medium

- Procedure

## The sampling point should be locked.

- Disinfect the sampling point with 70% IPA and flame the outlet where applicable

- The water should be allowed to flow for 2 minutes before sampling.

- The sterile sample bottle should be opened momentarily and sample taken up to ¾ of its volume in the presence of burning flame to achieve sterility of the neighborhood

- Do not allow the ‘bottle’ or the ‘water in the bottle’ to come in contact with the valve

- Remove the bottle from the sample stream and place the cap on the bottle as quickly as possible

- Close the valve

- The sample should be taken to the laboratory and labeled with the sample point

For Sterility testing

- Ensure the sterile room is cleaned, the workbench and the laminar flow is disinfected with 70% IPA.

- Prepare Tryptone soy broth for fungal growth and fluid Thioglycollate medium for bacterial growth as per sop for media preparation.

- Dispense 100ml of each broth into sterile bottles and sterilize for 15 minutes at 1210C, allow cooling to about 450C

- Get your WFI water sample inside a sterile container

- Label each sample bottle appropriately.

- Dispense 20ml of the water sample and inoculate 10ml into the broth (FTM and TSB) respectively.

- Incubate the inoculated samples at 20-250C for 72hours (TSB) and 30-350C for 48hours (FTM) for seven days (7).

Observation

- Observe any growth in the sample for the 7 days of incubation

CONCLUSION:

- The water sample is said to be sterile if no growth (cloudiness) is observed after 7th day of incubation.

REPORTING:

- The result should be recorded in water analysis record (BML-QDF-MC-014-F01)

- The final result should be reported to QAM

- If the specification is exceeded, there is need for re-sanitization and this should be reported to the appropriate supt. Pharmacist.

Materials, Equipment, Apparatus and Reagents

- Sterile petri-dishes

- Syringe

- Cotton wool

- 70% IPA

- Aluminum foil

- Lamina flow cabinet

- Bunsen burner

- Spatula

- Incubator

- Adhesive Tape

- Media. i.e. TSB,FTM

6.0 REFERENCE/ENCLOSURE

REFERENCE: USP <1234> Water for pharmaceutical purposes

USP <71> Sterility Test

ENCLOSURE 1: Water Analysis Record-(BML-QDF-MC-014-F01)

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |