- OBJECTIVE:

To provide written guidelines for calibration of Electronic weighing balance to ensure accurate and reliable measurement during sampling operations.

- SCOPE:

This SOP covers the process of calibration of Electronic weighing balance.

3.0 RESPONSIBILITY/ACCOUNTABILITY:

Laboratory Attendant, Quality Control Microbiologist, Quality control Head, Quality Assurance Manager, Superintendent Pharmacist.

4.0 ABBREVIATIONS AND DEFINITIONS:

N/A

- Procedure

- Place the weighing balance on a stable, horizontal surface.

- Ensure the weighing balance is clean and free from dust

- Ensure the weighing pan is empty

Calibration Steps

- Press  the ON/OFF key to switch on the balance

- Allow the display to stabilize and show zero (0.00)

- Press and hold the MODE key until display shows “CAL”

- When the display flashes “200.0”, it indicates the required 200g calibration weight

- Carefully place the 200g standard weight at the center of the weighing pan

- Wait for approximately 3-4 seconds for the reading to stabilize

- Once the reading stabilizes, the balance will complete the calibration automatically.

- Remove the calibration weight from the weighing pan

- Verify the return against 20g, 50g and 100g

- Press the ON/OFF key to switch off the balance

- Restart the balance to confirm proper operation

- Record the reading in the calibration logbook (BML.QDF.MC.045-F01)

Safety Precaution

- Avoid exposure to extreme heat or cold

- Place the weight gently at the center of the weighing pan

- Use certified standard calibration weights only

- Handle calibration weight carefully to prevent contamination or damage

Materials, Equipment, Apparatus and Reagents

- Electronic weighing balance

- Standard Calibration Weight

- Calibration Logbook

6.0 REFERENCE/ENCLOSURE:

ENCLOSURE 1- ELECTRONIC WEIGHING BALANCE CALIBRATION LOGBOOK- (BML.QDF.MC.045-F01).

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY:

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |