- OBJECTIVE:

To describe the procedure of Calibration of Conductivity Meter.

- SCOPE:

To provide written procedure for calibration of Conductivity Meter

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Quality Control Analyst: must follow the procedure as per the SOP for calibration of the conductivity meter and perform documentation related to the operation, calibration and maintenance in the respective log.

- Quality Control Head: must train all the concerned QC Analyst before the implementation of the SOP, ensure the operation and calibration of the instrument as per SOP and ensure proper documentation

- Quality Assurance Manager approve the SOP

- ABBREVIATION AND DEFINITIONS

- GLP: Good Laboratory Practice

- GDP: Good Documentation Practice

DEFINITION

- Conductivity: it is the ability of a solution to conduct electricity and the electricity is conducted in solution by ions present unit of measurement of conductivity is Siemen(S)

- PROCEDURES

- Connect the meter to the power supply and switch on.

- Preheat the meter for 30 minutes before calibration.

- Rinse the conductivity probe with distilled water and then with a small portion of the standard solution.

- Pour fresh 0.01 M KCl standard solution into a clean beaker.

- Immerse the probe into the standard solution ensuring no air bubbles and proper immersion depth.

- Allow the reading to stabilize and record the observed conductivity value (Kc).

Calibration Adjustment:

- The standard conductivity value (K) for 0.01 M KCl is:

1413 µS/cm at 25°C

- Calculate the cell constant (J) using:

J = K / Kc

Where:

K = standard conductivity (1413 µS/cm at 25°C)

Kc = observed conductivity

- Set the cell constant on the meter:

- Press keys sequentially:

SETUP → UP → SETUP → OK

- Input calculated value (J), then press OK and EXIT.

Temperature Consideration:

- Ensure calibration is carried out at 25°C, or use automatic temperature compensation if available.

- If temperature differs from 25°C, apply appropriate correction or allow solution to equilibrate.

Acceptance Criteria:

- The calibrated reading must be within ±2% of the standard value.

Frequency:

- Calibrate daily before use.

- Recalibrate when drift is observed or after electrode cleaning.

Materials, Equipment, Apparatus and Reagents

- 0.01 M KCl standard solution

- Distilled water

- Clean beakers

Safety, Health and Environmental Precautions

- Use appropriate PPE (lab coat, gloves).

- REFERENCE AND ENCLOSURE

- Equipment Manual

- ENCLOSURE 1- CONDUCTIVITY METER CALIBRATION LOG - BML-QDG-QA-036-F01

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | All Quality Control Analyst |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |